|  | СООМЕТ-21/12.1 (2) |
| --- | --- |

|  | ##### COOMET Recommendation | CОOMET R/GM/__:20__ |
| --- | --- | --- |
|  | Use of concepts “error of measurement” and “uncertainty of measurement”. General principles materials |  |
| Approved in __the meeting of COOMET Committee ( , 20__) |  |  |

1. Introduction

In this document concepts «error of measurement» and «uncertainty of measurement» are analyzed, and recommendations about logically consistent joint application of these concepts of various metrological problems are made.  As the basic terminological documents, [1 – 3] are accepted.

These recommendations are developed on the issue of CООМЕТ 347/RU/05.

The recommendations stated in this document are intended for use in legislative and applied metrology.

Terminological articles in these recommendations, repeating terminological articles of [2], are represented by thin lines.

2. Scope

These recommendations are intended for use by working out of standard documents of all kinds, scientific and technical ones, educational and reference books on metrology, works entering into the sphere on standardization and (or) these works results use.

3. Initial provisions

In [1] basic distinction of concepts «error of measurement» and «uncertainty of measurement» is underlined, but possibility use of concept «error» is not excluded. Thus, it is meant that the concrete error always has a certain sign (positive or negative).  Application of concept «true value of the measured quantity» instead of concept «reference quantity value » in error definition in [2] it does not change its sense. By definition, unlike «error» concept, «uncertainty» characterizes dispersion of values which could be attributed to the measured quantity.

The incorrectness of application of concept «error» is indicated in  its mixing with the concepts  different by implication, such as « measurement result error characteristics », « error confident limits». The error of a certain result of measurement is evinced in the considered experiment with a certain copy of a measuring instrument, and during the evaluation of  «error characteristics» a set of possible values of errors in virtual or real experiments with various copies of measuring instruments of the given type at admissible variation of measurement conditions is used. Therefore the standard evaluation of the standard deviation, non-excluded systematic error and confident limits of a set of errors of measurement results do not correspond any more to initial definition of an error. Actually, these evaluations characterize not an error, but a dispersion of values attributed to the measured quantity on the basis of the used information, i.e. uncertainty.

Concepts «error of measurement» and «uncertainty of measurement» should be applied according to their definitions, without error substitution by evaluation parameter and components of measurement results dispersion.

4. Basic terms, definitions and comments to them

In these recommendations, the following terms with appropriate definitions are used:

4.1

measurement error (error of measurement, error): measured quantity value minus a reference quantity value N o t e 1: The concept of 'measurement error' can be used both a) when there is a single reference quantity value to refer to, which occurs if a calibration is made by means of a measurement standard with a measured quantity value having a negligible measurement uncertainty, or if a conventional quantity value is given, in which case the measurement error is known, and b) if a measurand is supposed to be represented by a unique true quantity value or a set of true quantity values of negligible range, which case the measurement error is not known. N o t e 2: Measurement error should not be confused with production error or mistake. [2, article 2.16]

C o m m e n t: True value of a quantity cannot be defined. This concept applies only in theoretical researches. In practice, they use reference quantity value Xо (see 4.2), and a error of measurement Δ is defined according to formula:

Δ = Xm – Xо ,

where Xm – the value of quantity received by measurement (result of measurement [2, article 2.9]);

Xо – the value attributed to concrete quantity and accepted, often under the agreement, as a value, which has an uncertainty acceptable for the given purpose [2, article 2.12].

Thus, by definition, concept «the measurement error» concerns only the certain result of  measurement received with the use of a certain copy of a measuring instrument. The error of  measurement is a certain positive or a negative number. There are no bases to assign to this concept a sense of statistical parameter of any set of real or guess values. «Error of measurement» and «uncertainty of measurement»  represent different concepts; they should not be confused with each other or used in the wrong way [1, point 3.2.2].

4.2.

reference quantity value (reference value): quantity value used as a basis for comparison with values of quantities of the same kind N o t e 1: A reference quantity value can be a true quantity value of a measurand, in which case it is unknown, or a conventional quantity value, in which case it is known. N o t e 2: A reference quantity value with associated measurement uncertainty is usually provided with reference to a) a material, e.g. certified reference material, b) a device, e.g. stabilized laser, c) a reference measurement procedure, d) a comparison of measurement standards. [2, article 5.18]

C o m m e n t:  Concept «conventional quantity value» [2, article 2.12] covers the concept «the real value of quantity » («the conventional true quantity value») ─ the value of quantity received experimentally and being so close to the true value, that in the put measuring problem it can be used instead of that.

4.3. relative error: Ratio Δ/Xo error of  measurement Δ to reference quantity value Xо.

C o m m e n t: It is not recommended to replace in this respect the  reference quantity value by the result of measurement, as it contradicts the concept of « error of measurement » definition.

4.4.

systematic measurement error (systematic error of measurement, systematic error): component of measurement error that in replicate measurements remains constant or varies in a predictable manner N o t e 1: A reference quantity value for a systematic measurement error is a true quantity value, or a measured quantity value of a measurement standard of negligible measurement uncertainty, or a conventional quantity value. N o t e 2: Systematic measurement error, and its causes, can be known or unknown. A correction can be applied to compensate for a known systematic measurement error. N o t e 3: Systematic measurement error equals measurement error minus random measurement error. [2, article 2.17]

C o m m e n t: It is necessary to keep in mind that at definition of a difference of the specified errors, each of the errors is  taken with the positive or negative sign.

4.5.

random measurement error (random error of measurement, random error): component of measurement error that in repeated measurements varies in an unpredictable manner N o t e 1: A reference quantity value for random measurement error is the average that would ensue from an infinite number of replicate measurements of the same measurand. N o t e 2: Random measurement errors of a set of replicate measurements form a distribution that can be summarized by its expectation, which is generally assumed to be zero, and its variance. N o t e 3: Random measurement error equals measurement error minus systematic measurement error. [2, article 2.19]

C o m m e n t: It is necessary to keep in mind that at definition of a difference of the specified errors, each of the errors is taken with the positive or negative sign.

4.6

measurement uncertainty (uncertainty of measurement, uncertainty): Non-negative parameter characterizing the dispersion of the quantity values attributed to a measurand, basis on the information used N o t e 1: Measurement uncertainty includes components arising from systematic effects, such as components associated with corrections and the assigned quantity values of measurement standards, as well as the definitional uncertainty. Sometimes estimated systematic effects are not corrected for but, instead, associated measurement uncertainty components are incorporated. N o t e 2: The parameter may be, for example, a standard deviation called standard measurement uncertainty (or a specified multiple of it), or the half-width of an interval, having a stated coverage probability. N o t e 3: Measurement uncertainty comprises, in general, many components. Some of these may be evaluated by Type A evaluation of measurement uncertainty from the statistical distribution of the quantity values from series of measurements, and can be characterized by standard deviations. The other components, which may be evaluated by Type B evaluation of measurement uncertainty, can also be characterized by standard deviations, evaluated from probability density functions based on experience or other information. N o t e 4: In general, for a given set of information, it is understood that the measurement uncertainty is associated with a stated quantity value attributed to the measurand. A modification of this value results in a modification of the associated uncertainty [2, article 2.26]

C o m m e n t: The value said in the note 4 is the best estimation of value of the measurand, and all components of uncertainty, including the components caused by systematic effects, for example, connected with corrections and standards, lead to dispersion [1, point 2.2.3, Notes 1-3].

Quantitatively, «uncertainty of measurement» (as a rule) can be characterized by «standard uncertainty of measurement» – uncertainty of result of the measurement expressed as a standard deviation, or «the expanded measurement uncertainty» [1, points 2.3.1, 2.3.4 and 2.3.5].

Thus, «uncertainty of measurement» as the parameter, characterizes the dispersion of a set of measurement results possible values in a considered measuring situation, but not an error of concrete result of measurement. For example, a case is possible, when the result of measurement has negligibly small error with large uncertainty [1, point 3.3.1, the note].

5. Recommendations concerning correct application of concepts «error of measurement» and «uncertainty of measurement»

5.1 Application of concepts «error of measurement» and «uncertainty of measurement» in concrete metrological situations

5.1.1 Result of measurement is the value of quantity received by its measurement and uncertainty of measurement . Concrete results of measurements in any metrological situations unequivocally can and should be characterized by uncertainty. Application of concept error of measurement result, which is essentially unknown and is specifically indefinable, is possible only in theoretical reasoning concerning measurement results. Error concept is supposed to be used during measuring instrument calibration and verification (see point 4.1, note 1, enumeration a).

5.1.2 Results of measurements, which are carried out during national standards comparisons (key, regional, international) according to the Arrangement [4], are represented with the detailed data about evaluation of uncertainty. Here, the errors characteristics of standards, which are stated in registration certificates for national standards, can be used for evaluation of uncertainty of measurement result.

5.1.3 Calibration and measuring capabilities of national metrological institutes, according to Appendix C of Agreement [4], are represented with indication expanded uncertainty of measurement results and coverage factor. Here, it is obligatory to indicate data about metrological traceability under transfer of appropriate measurement scale or size of measurement unit.

5.1.4 In measurement procedure, operation and calculus complex is described; its fulfillment provides obtaining of the measurement result with the established indices of accuracy.

Results of measurement according to measurement procedure are recommended to be accompanied by uncertainty of measurement evaluations. Measurement procedure may include information about target uncertainty of measurement [2, point 2.34]. Standardized (certified) measurement procedure may maintain other established indices of accuracy measurements, for example by [5], corresponding of concrete setting and aria of its application.

5.1.5 At calibration of measuring instruments, they establish, under certain conditions, the proportion between values of quantity according to indications of measuring instrument and appropriate values realized using the standard. According to calibration results, corrections to measuring instruments indications may be made or values realized by measuring instruments  may be made more exact. As it is specified in Appendix C to Agreement [4], only uncertainty can be a calibration accuracy characteristic. The same concerns the results of measuring instruments graduation in the course of calibration.

5.1.6 Measuring instruments metrological characteristics are standardized using the concept «error» or «accuracy class». Here, permissible error limits for measuring instruments of a given type are used.

5.1.7  Measuring instruments verification consists in establishing of measuring instrument (MI) suitability to application on the basis of experimental definition of metrological characteristics and acknowledgement of their conformity specified requirements. At verification, measurement standards are used.

MI verification may consist in the following:

a) Determination of MI suitability to application with rejection of those MI, whose error exceeds limits of the assumed error, established for MI of a given type;

b) Establishing of conventional true values or gauging characteristics of the MI brought for verification (including the way of corrections application);

c) Determination of MI suitability to application according to norms of their stability (with rejection of those MI, whose change of the conventional true value or gauging characteristics exceeded permissible instability limit established for MI of a given type for inter-verification interval), and establishing of the conventional true values or gauging characteristics of the other MIs.

At verification, norms of measuring instruments error limits are operated. Therefore, in verification  it is assumed to specify, in what proportion should be the expanded uncertainty of measurements during verification and limits of errors of measuring instruments of the given confirmed type, and also to specify criteria of measuring instruments validity taking into account uncertainty of measurements at verification (see Appendix A). Here, general principles taking into account  uncertainty of measurements in procedure  evaluation of conformity with specified requirements may be used  [6].

5.1.8 In construction of a verification chain, accuracy of unit size transfer method may be characterized by uncertainty of measurement.

5.2 General recommendation.

On the basis of the considered metrological situations, it is possible to assume the general rule: in the most metrological situations, measurement results are characterized by uncertainty, and accuracy of measuring instruments are characterized by error limits. Concept «error» is used at comparison with  reference quantity value, and evaluation errors are obtained at calibration or verification of measuring instruments.  Thus, concepts "uncertainty" and "error" are recommended to be used harmoniously, without mutual contraposition and exclusion of one of them.

Appendix A

(reference)

Explanations to the use of concept «uncertainty of measurement» during verification

In the international document OIML D 8 [7], the concept «uncertainty of measurement during verification of measuring instruments» is used without this provision concretization. Different versions of its realization are possible due to the diversity of real situations in correlation of error limits of verified measuring instruments and standards used for that, in correlation of type A and type B evaluated uncertainties. For example, if the expanded uncertainty of measurements at verification (coverage factor 2) does not exceed 1/3 of limits of error, uncertainty is ignored. The other possible versions of measuring instruments (MI) suitability criteria on verification results are:  MI evaluated error doesn’t exceed difference limit of error and expanded measurement uncertainty during verification (see, for example  [8]); MI evaluated error does not exceed square root from difference of squares limit of error and expanded measurement uncertainty during verification. Such criteria of suitability may be used, in particular, during verification of MI representing the measures, for which the error is «difference between nominal values of measures and real values of quantities reproduced by them». The sense of uncertainty account during verification is explained by the graph in Figure А.1. The versions represented in this Appendix explain only the principles of concept «error» and «uncertainty» use, but they are not preferably recommended and don’t cover all the diversity of possible situations.

U

U

X

Xи

Xд

Xи+U

Xи-U

Xд-Δн

Xд+ Δв

Хд – reference (real) value (of standard);

Хи –indication of tested measuring instrument;

Dн, Dв – upper and lower limits of error, according to normative document for measuring instruments to be tested (usually  Dн = Dв = D);

U – expanded uncertainty.

Figure А.1 – Diagram of uncertainty taking into account at confirmation of conformity of tested measuring instrument to limits of error

Bibliography

| [1] ISO/IEC Guide 98-3:2008 | Uncertainty of measurement – Part 3: Guide to the expression of uncertainty in measurement (GUM:1995) |
| --- | --- |
| [2] JCGM 200:2008 (E/F) | International vocabulary of metrology – Basic and general concepts and associated terms (VIM). |
| [3] РМГ 91-2009 | Государственная система обеспечения единства измерений. Совместное использование понятий «погрешность измерения» и «неопределенность измерения». Общие принципы. (State system for ensuring the uniformity of measurements. Joint use of concepts «error of measurement» and «uncertainty of measurement». General principles.) |
| [4] Mutual recognition of national measurement standards and of calibration and measurement certificates issued by national metrology institutes , Paris, 14 October 1999. |  |
| [5] ISO 5725-1:1994 | - Accuracy (trueness and precision) of measurement methods and results. Part 1: General principles and definitions |
| [6] ISO 10576-1:2003 | Statistical methods. Guidelines for the evaluation of conformity with specified requirements. Part 1. General principles. |
| [7] International Document OIML D 8, Edition 2004 | Measurement standards. Choice, recognition, use, conservation and documentation. |
| [8] OIML R 111-1-2004 | Weights of classes E1, E2, F1, F2, M1, M1-2, М2, М2-3 и M3. Part 1: Metrological and technical requirements. |

Information data

1. Coordinating organization: «NATIONAL Scientific Research Institute for PhysicOTechnical and radio ENGINEERING Measurements».

Authors: A.S.Doynikov (supervisor), B.N.Krupin, S.G.Bessonov (text English editors)

2. COOMET theme 347/RU/05.

3. Document was approved at the  __th meeting of COOMET Committee.

4. Information about application of the document in organizations – members of COOMET.

Recommendation is developed on the basis of РМГ 91-2009 « National system of uniformity of measurements. Uniformity of “measurement error” and “measure ment uncertainty” concepts joint use. General principles.» and of materials of publications:

A.S. Doynikov. Methodical recommendations about application of concepts "error" and "uncertainty" in various metrological tasks. - “Legislative and applied metrology”, 2006, №1, with. 43-46;

A.S. Doynikov “Consistent application of concepts “error” and “uncertainty of measurements” – Metrology and metrological maintenance.  Theses of reports of the International scientific and technical conference. ∕ edited by  V. N. Koreshkova, et al. – Minsk: BelGIM, 2007, p. 80-84.

Recommended conceptually-terminological system and general principles are intended for use in legislative and applied metrology.

Contents

1. Introduction . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . .       1

2. Scope . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .       1

3. Initial provisions. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .       1

4. Basic terms, definitions and comments to them. . . . . . . . . . . . . . . . . . . . .       2

5. Recommendations concerning correct application of concepts

«a measurement error» and «uncertainty of measurement» . . . . . . . . . . .            5

Appendix A (reference). Explanations to the use of concept «uncertainty

of measurement» during verification . . . . . . . . . . . . . . . . . . . . . . . . . . . .        7

Bibliography. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 8

Information data . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .9